Spectral Fitting

for Monophosphonated Trityl Radicals

User Guide

Driesschaert Lab

In Vivo Multifunctional Magnetic Resonance Center, Robert C. Byrd Health Sciences Center

Department of Pharmaceutical Sciences, School of Pharmacy

West Virginia University, Morgantown, WV, 26506, USA

Table of Contents

Background Information about Monophosphonated Trityl EPR	3

Definitions	4

Fitting Parameters	4

Calibration and Results	4

Installation	5

Option 1 – standalone executable	5

Option 2 – installation	5

Spectra Fitting	6

Selecting a Spectrum File	6

Fitting a Spectrum with an Existing Calibration	7

Editing a Calibration	9

Calibration	10

Starting a New Calibration	10

Determining the Fitting Parameters	10

Acidic Peak (Outer Peaks) Linewidths	10

Basic Peak (Inner Peaks) Linewidths	11

A/B Distance	12

Alternate - Both Peaks Linewidths (plus A/B Distance)	13

Calibrating Slopes	14

Oxygen Slope of Acidic Peak (Outer Peaks)	14

Oxygen Slope of Basic Peak (Inner Peaks)	15

Alternate - Oxygen Slope of Acidic/Basic Peaks at pH 7	16

Phosphate Slope and Intercept	17

pKa	18

Saving the New Calibration	18

Additional Notes	19

Saved Calibration File	19

# Background Information about Monophosphonated Trityl EPR

Example monophosphonated trityl (pTAM/HOPE) full EPR signal using L-band EPR spectroscopy.

lowfield component A/B distance < 0

Basic peak

Acidic peak

highfield component A/B distance > 0

Acidic peak

Basic peak

*Note – all instructions included in this guide utilize half spectra (highfield/lowfield components) of the monophosphonated trityl EPR signal. Full signal spectra may also be used by adjusting the center offset to align with the component of interest.

# Definitions

## Fitting Parameters

- Gaussian LW (milliGauss) – Gaussian contribution to line shape for both the acidic and basic peaks

- Lorentzian A LW (milliGauss) – Lorentzian contribution to line shape for the acidic peak with no oxygen

- Lorentzian B LW (milliGauss) – Lorentzian contribution to line shape for the basic peak with no oxygen

- A/B Distance (milliGauss) – Distance between acidic and basic peaks in a buffer free solution

- Oxygen LW (milliGauss) – Lorentzian linewidth broadening due to oxygen

- Acidic Fraction – fraction of pTAM in the acidic state (H-pTAM3-)

- H+ Exchange (1/s) – proton exchange rate (ka)

- Pi Term  - adjustment to the H+ exchange rate based on pH and Ka of the buffer (phosphate), linear relationship with phosphate concentration

Pi Term= kaH++KaBuf2πKaBuf

- Center Offset (points) – adjustment to the center field position

- Phase Correction (degrees) – adjustment to the phase. To flip, enter 180.

- Start Value – Center of the fitting range for the parameter of interest. If the parameter switch is set to Constant and/or the Range is 0, the Start Value is the assigned value for that parameter.

- Range – Defines how much the parameter can vary when fitting in the positive/negative direction of the Start Value. (Example – If Start value = 100, Range = 50, the fitting range will be 50-150). The Range cannot be greater than the absolute value of the Start Value because the fitting parameters cannot below zero (except A/B Distance is negative when looking at the low-field component).

- Best Fit – The result for the corresponding parameter for the best fit of the spectrum.

## Calibration and Results

- Oxygen (pO2)

  - Slope – The linear relationship between oxygen-induced line broadening (Oxygen LW, mG) and the partial pressure of oxygen (mmHg).

  - Result – The oxygen partial pressure derived from the Oxygen LW of the fitted spectrum and the Oxygen slope.

- pH

  - pKa – The pKa of pTAM between the acidic and basic states : pTAM3- ↔ pTAM4-

  - Result – The pH derived from the Acidic Fraction of the fitted spectrum and the pKa.

- Phosphate

  - Slope – The linear relationship between the Pi Term (1/s) and the inorganic phosphate concentration [Pi] (mM).

  - Result – The inorganic phosphate concentration (mM) derived from the H+ Exchange and Pi Term of the fitted spectrum and the phosphate slope.

# Installation

You have 2 options for “installing” this software on your PC :

- Using the executable in standalone form (requires MatLab Runtime)

- Run the installation file pTAMAppInstaller_mrc.exe (includes MatLab Runtime).

Disclaimer – this software was developed and tested in a Windows 10 environment and may not run as expected on Macs and Linux operating systems.

## Option 1 – standalone executable

- Unzip the file containing the application files, pTAM_Fitting_App.exe and pTAM_cal.txt.

- Store these files together wherever you would like to tun the program from.

  - This will perform best if in a file that has writable permissions.

- To run the executable, you will need to install MatLab Runtime (if not already installed) which can be found here:

https://www.mathworks.com/products/compiler/matlab-runtime.html

## Option 2 – installation

- Run pTAMAppInstaller_mrc.exe (includes the MatLab Runtime) or pTAMAppInstaller_web (downloads Matlab Runtime).

- The first screen provides a summary of the software. Click Next.

- The second screen allows you to select the installation folder. You can also choose to add a shortcut. Choose and Click Next. (Defaults to “C:\Program Files\pTAM_Fitting_App”)

- The third screen checks for Matlab Runtime. Review and click Next.

- The fourth screen is a confirmation page. Review and click Install.

- Installation is complete when the installer reads “Installation completed successfully.”

  - The original pTAM_cal.txt file for saving calibrations will be installed under the same file as the application.  Because the Program Files folders generally do not have writable permissions, a writeable copy will be made to your user directory (C:\Users\[username]\pTAM_Fitting_App). If a user directory is not available, a writeable copy will be made to the deployable archive (generally Temp app data).

# Spectra Fitting

## Selecting a Spectrum File

- Click  Select File  in the top left box.

- Navigate to the location on your computer of your spectrum file and select the file of interest.

*Note - The software currently supports “.DSC”, “.DTA” (Bruker), or “.SPE” (Magnettech) files.

- Click  Open .

- The spectrum preview will be displayed in the graph below with a blue line.

- Choose whether the spectrum is in Gauss (G) or millitesla (mT) before fitting the spectrum. (DSC and DTA files will default to G; SPE files will default to mT) See the next section (page 7) for fitting.

*Note – Once in the correct file path/directory, you may type changes to the Data file name to easily switch between files in the same path.

## Fitting a Spectrum with an Existing Calibration

- Choose the saved calibration from the drop-down menu at the top right of the application.

- Ensure that you have a spectrum file selected (See page 6).

- Choose whether the spectrum is in Gauss (G) or millitesla (mT), if not already done.

- Ensure that the Oxygen LW, Acidic Fraction, and H+ Fraction parameters are set to vary an appropriate amount.

*Note – The varying Range cannot be greater than the absolute value of Start Value (parameters cannot be negative).

- Click  Fit Spectrum .

- The fitted spectrum (red line) will appear in the graph with the experimental data (blue line). The residual difference between the experimental and fitted spectra will appear in the graph below.

- The Best Fit for the parameters Oxygen LW, Acidic Fraction, and H+ Fraction (and the related Pi Term) will appear to the right.

*Note – The Pi term = kaH++KaB2πKaB (Hz) and has a linear relationship with the inorganic phosphate concentration.

- The resulting Oxygen Partial Pressure (mmHg), pH, and Inorganic Phosphate Concentration (mM) will be calculated based on the slopes and pKa shown, and they will be displayed at the bottom right.

- If the spectrum does not successfully fit due offset positioning, adjust the Center Offset to move the simulated spectrum left (negative) or right (positive) and re-click  Fit Spectrum . Intervals of 25 are recommended.

- If the spectrum does not successfully fit due to the phase, enter a guessed degree adjustment into Phase Correction (example=180 to flip) and re-click  Fit Spectrum .

*Note - Both center and phase do have some variance built into the fitting algorithm. These adjustments should only be needed for more exaggerated offset or out of phase spectra.

## Editing a Calibration

  - Select the calibration you want to work with in the drop-down menu at the top right of the application.

  - Click  Edit Cal  below the drop down menu.  This will allow you to change the fitting parameters, calibrated slopes, and pKa. The following values are saved:

Gaussian LW, Lorentzian A LW, Lorentzian B LW, A/B Distance,

Oxygen slope, pKa, Phosphate slope

  - You can follow the section on pages 7-8 to fit a spectrum with your edited calibration.

  - If you would like to save these changes to the current calibration, Click  Save Changes . A dialog box will appear and ask if you would like to change the name of the calibration (It will default to the current name). Make any changes the name or leave it the same and click  OK .  Hitting Cancel or submitting a blank name will cancel the save.

  - If you would like to save the changes as a new calibration, Click  Save As . A dialog box will appear asking for the new calibration name. Enter the name you would like and Click  OK . Hitting Cancel or submitting a blank name will cancel the save.

  - If you would like to delete this calibration, Click  Delete Cal . A prompt will ask you if you are sure you would like to delete this calibration. If Yes, the calibration will be deleted. If No, the deletion will be cancelled.

# Calibration

## Starting a New Calibration

- To make a new calibration, select “0 Custom/Calibrate” the drop-down menu at the top right of the application.

- A panel of shortcut buttons for calibration will appear below the drop-down menu.

## Determining the Fitting Parameters

### Acidic Peak (Outer Peaks) Linewidths

**EPR Sample preparation and spectrum acquisition**

Prepare an acidic sample of suitable pTAM concentration and volume for your EPR instrument. Include NaCl and phosphate buffer as appropriate. The pH should be approximately between 4 and 5. Remove all oxygen from the sample before recording the EPR signal via a gas controller or the addition of glucose oxidase/glucose (the latter method may affect the pH).

- With “0 Custom/Calibrate” selected and the “Quick Cal Buttons” panel showing, click  pH 5 LW . This automatically sets the Gaussian and acidic peak Lorentzian (Lorentzian A) linewidths to vary. It also holds the oxygen-induced line broadening (Oxygen LW) at 0 and the Acidic Fraction at 1.  Lorentzian B LW, A/B Distance, and H+ Exchange can be ignored with just the acidic peak present.

- Select a file (see page 6) of a spectrum recorded at pH~5 and with 0% oxygen. Click  Fit Spectrum .

- Record the Best Fit results for Gaussian LW and Lorentzian A LW.

- Repeat for all replicate spectra. Average the results for the two parameters and enter the averages into the corresponding Start Values.

### Basic Peak (Inner Peaks) Linewidths

**EPR Sample preparation and spectrum acquisition**

Prepare a basic sample of suitable pTAM concentration and volume for the EPR instrument. Include NaCl and phosphate buffer as appropriate. The pH should be between 10 and 12. Remove all oxygen from the sample before recording the EPR signal via a gas controller or the addition of glucose oxidase/glucose (the latter method may affect the pH).

- With “0 Custom/Calibrate” selected and the “Quick Cal Buttons” panel showing, click  pH 10+ LW . This automatically sets the overall Gaussian and basic peak Lorentzian (Lorentzian B) linewidths to vary. It also holds the oxygen-induced line broadening (Oxygen LW) at 0 and the Acidic Fraction at 0.  Lorentzian A LW, A/B Distance, and H+ Exchange can be with just the basic peak present.

- Select a file (see page 6) of a spectrum recorded at pH>10 and with 0% oxygen. Click  Fit Spectrum .

- Record the Best Fit for Gaussian LW and Lorentzian B LW.

- Repeat for all replicate spectra. Average the results for the two parameters and enter the averages into the corresponding Start Values.

- Once you have fit both the acidic and basic peaks, average the Gaussian LW for both together for the final value. Enter that for the Start Value.

### A/B Distance

**EPR Sample preparation and spectrum acquisition**

Prepare a sample of suitable pTAM concentration and volume for the EPR instrument. Include NaCal as appropriate, but do not include phosphate buffer for determining the A/B Distance. The pH should be approximately 7. Remove all oxygen from the sample before recording the EPR signal via a gas controller.

- With “0 Custom/Calibrate” selected and the “Quick Cal Buttons” panel showing, click  0 Pi Dist . This automatically sets the overall Gaussian, acidic peak Lorentzian (Lorentzian A), and basic peak Lorentzian (Lorentzian B) linewidths to constant. Ensure they are the values you determined for this calibration (pages 10 and 11). It also holds the oxygen-induced line broadening (Oxygen LW) at 0 and H+ Exchange at 1e-5.  It allows the acidic fraction and A/B Distance to vary. A/B Distance Start Value should be positive for the high-field/right component (acidic peak on right) and negative for the low-field/left component (acidic on left).

- Select a file (see page 6) of a spectrum recorded at pH 7 with no phosphate and with 0% oxygen.

- Click  Fit Spectrum .

- Record the Best Fit for A/B Distance.

- Repeat for all replicate spectra. Average the results for the parameter and enter the average into the corresponding Start Value.

### Alternate - Both Peaks Linewidths (plus A/B Distance)

**EPR Sample preparation and spectrum acquisition**

Prepare a sample of suitable pTAM concentration and volume for the EPR instrument. Include NaCal as appropriate, but do not include phosphate buffer for determining the A/B Distance. The pH should be approximately 7. Remove all oxygen from the sample before recording the EPR signal via a gas controller.

- With “0 Custom/Calibrate” selected and the “Quick Cal Buttons” panel showing, click  pH 7 LW . This automatically sets the overall Gaussian, acidic peak Lorentzian (Lorentzian A), and basic peak Lorentzian (Lorentzian B) linewidths to vary. It also allows the acidic fraction and A/B Distance to vary.  It holds the oxygen-induced line broadening (Oxygen LW) at 0 and H+ Exchange at 1e-5.  A/B Distance Start Value should be positive for the high-field/right component (acidic peak on right) and negative for the low-field/left component (acidic on left).

- Select a file (see page 6) of a spectrum recorded at pH 7 with no phosphate and with 0% oxygen.

- Click  Fit Spectrum .

- Record the Best Fit for Gaussian LW, Lorent A LW, Lorent B LW, and A/B Distance.

- Repeat for all replicate spectra. Average the results for the four parameters and enter the averages into the corresponding Start Values.

## Calibrating Slopes

### Oxygen Slope of Acidic Peak (Outer Peaks)

**EPR Sample preparation and spectrum acquisition**

Prepare an acidic sample of suitable pTAM concentration and volume for the EPR instrument. Include NaCl and phosphate buffer as appropriate. The pH should be approximately between 4 and 5. This should be the same sample as page 10. Adjust the oxygen partial pressure surrounding or bubbled through the sample with a gas composition controller. Record the spectra at varying percentages of oxygen below air (21%).

- Click  pH 5 O2 Cal .  This automatically sets the oxygen-induced line broadening (Oxygen LW) to vary. It also holds the Gaussian LW and acidic peak Lorentzian (Lorentzian A LW) linewidths constant and the acidic fraction at 1.  Ensure these are the values you determined for this calibration (page 10). Lorentzian B, A/B Distance, and H+ Exchange can be ignored with just the acidic peak present.

- Select a file (see page 6) of a spectrum recorded at pH~5 with a known percentage of oxygen.

- Click  Fit Spectrum . Record the Best Fit for Oxygen LW.

- Repeat for all replicate spectra.

- Repeat for each known percentage of oxygen

- Plot the Oxygen LW (y; mG) against oxygen partial pressure (x; mmHg) to determine the linear relationship/slope.

- Once you find the acidic peak oxygen slope and the basic peak oxygen slope, average them together and enter the value for the Oxygen (pO2) slope in the Calibrations and Results section.

### Oxygen Slope of Basic Peak (Inner Peaks)

**EPR Sample preparation and spectrum acquisition**

Prepare a basic sample of appropriate pTAM concentration and volume for the EPR instrument. Include NaCl and phosphate buffer as appropriate. The pH should be between 10 and 12. This should be the same sample as page 11. Adjust the oxygen partial pressure surrounding or bubbled through the sample with a gas composition controller. Record the spectra at varying percentages of oxygen below air (21%).

- Click  pH 10+ O2 Cal .  This automatically sets the oxygen-induced line broadening (Oxygen LW) to vary. It also holds the Gaussian LW and basic peak Lorentzian (Lorentzian B LW) linewidths constant and the Acidic Fraction at 0.  Ensure these are the values you determined for this calibration (page 11). Lorentzian A, A/B Distance, and H+ Exchange can be ignored with just the basic peak present.

- Select a file (see page 6) of a spectrum recorded at pH>10 with a known percentage of oxygen.

- Click  Fit Spectrum . Record the Best Fit for Oxygen LW.

- Repeat for all replicate spectra.

- Repeat for each known percentage of oxygen.

- Plot the Oxygen LW (y; mG) against oxygen partial pressure (x; mmHg) to determine the linear relationship/slope.

- Once you find the acidic peak oxygen slope and the basic peak oxygen slope, average them together and enter the value for the Oxygen (pO2) slope in the Calibrations and Results section.

### Alternate - Oxygen Slope of Acidic/Basic Peaks at pH 7

**EPR Sample preparation and spectrum acquisition**

Prepare a basic sample of suitable pTAM concentration and volume for the EPR instrument. Include NaCl and phosphate buffer as appropriate. The pH should be between 10 and 12. This should be the same sample as page 11. Adjust the oxygen partial pressure surrounding or bubbled through the sample with a gas composition controller. Record the spectra at varying percentages of oxygen below air (21%).

- Click  pH 7 O2 Cal .  This automatically sets the oxygen-induced line broadening (Oxygen LW) to vary. It also holds the Gaussian LW, acidic peak Lorentzian (Lorentzian A LW), and basic peak Lorentzian (Lorentzian B LW) linewidths and the A/B Distance constant.  Ensure these are the values you determined for this calibration. It also holds the H+ Exchange at 1e-5 and allows the Acidic Fraction to vary.

- Select a file (see page 6) of a spectrum recorded at pH~7 with a known percentage of oxygen.

- Click  Fit Spectrum . Record the Best Fit for Oxygen LW.

- Repeat for all replicate spectra.

- Repeat for each known percentage of oxygen

- Plot the Oxygen LW (y; mG) against oxygen partial pressure (x; mmHg) to determine the linear relationship/slope.

- Once you find the oxygen slope, enter the value for the Oxygen (pO2) slope in the Calibrations and Results section.

### Phosphate Slope and Intercept

**EPR Sample preparation and spectrum acquisition**

Prepare several samples of suitable pTAM concentration and volume for the EPR instrument. Include NaCal as appropriate.  Include different concentrations of inorganic phosphate up to 10mM. The pH of all should be approximately 7. Remove all oxygen from the samples before recording the EPR spectra via a gas controller.

- Click  Pi Cal .  This automatically sets the Acidic Fraction and H+ Exchange to vary. It also holds the Gaussian LW, acidic peak Lorentzian (Lorentzian A LW), and basic peak Lorentzian (Lorentzian B LW) linewidths and the A/B Distance constant.  Ensure these are the values you determined for this calibration. It also holds the Oxygen LW constant at 0.

- Select a file (see page 6) of a spectrum recorded at pH~7 with no oxygen and a known concentration of inorganic phosphate [Pi].

- Click  Fit Spectrum . Record the Best Fit for Acidic Fraction, H+ Exchange, and Pi Term.

- Repeat for all replicate spectra.

- Repeat for each known concentration if inorganic phosphate.

- Plot the Pi Term (y; 1/s) against inorganic phosphate concentration (x; mM) to determine the linear relationship/slope.

- Once you find the equation for the line, enter the values for the Phosphate [Pi] slope (kf) in the Calibrations and Results section.

### pKa

**EPR Sample preparation and spectrum acquisition**

Prepare a solution of suitable pTAM concentration for the EPR instrument and volume for titration and accurate pH measurement. Include NaCal as appropriate and 1mM phosphate buffer. The pH should be between 6 and 8.  Bubble in nitrogen, record pH, and take a sample for EPR measurement. Control gas composition in sample to all nitrogen before recording the EPR spectra. Titrate the solution with weak solutions of NaOH or HCl to adjust the pH to 5 or more values between 6 and 8. Repeat anerobic EPR measurements for each pH.

Note – only necessary if pKa is not known or defined in literature.

- Click  pKa Cal .  This automatically sets the Acidic Fraction and H+ Exchange to vary. It also holds the Gaussian LW, acidic peak Lorentzian (Lorentzian A LW), and basic peak Lorentzian (Lorentzian B LW) linewidths and the A/B Distance constant.  Ensure these are the values you determined for this calibration. It should also hold the Oxygen LW constant at 0.

- Select a file (see page 6) of a spectrum recorded with no oxygen and a known pH between 6 and 8.

- Click  Fit Spectrum . Record the Best Fit for Acidic Fraction.

- Repeat for all replicate spectra.

- Repeat for each known pH.

- Plot the Acidic Fraction (y) against pH (x) to plot the titration curve.

- Fit the data using the equation to find pKa:

y=1/(1+10x-pKa)

- Once you find the pKa, enter the value for the pKa in the Calibrations and Results section.

## Saving the New Calibration

- Select “0 Custom/Calibrate” in the drop-down menu at the top right of the application, if not already done.

- Make sure all of your determined fitting parameters are entered into the Start Values and calibrated slopes/intercepts/pKa are entered into the Calibration and Results section under Correlation. The following values are saved:

Gaussian LW, Lorentzian A LW, Lorentzian B LW, A/B Distance,

Oxygen slope, pKa, Phosphate slope

- Click  Save New Cal .

- A dialog box will appear asking for the new calibration name. Enter the name you would like and       Click  OK . Hitting Cancel or submitting a blank name will cancel the save.

# Additional Notes

## Saved Calibration File

- The calibrations are saved in a text file pTAM_cal.txt, which should be in the same directory as the software’s exe file.  If you installed the application in a read-only folder such as “C:\Program Files\pTAM_Fitting_App\”, a writeable copy of pTAM_cal.txt will be created in your user directory (Example – “C:\Users\[username]\pTAM_Fitting_App\”).

- The software will still run without pTAM_cal.txt, but will not load any existing calibrations and will only provide the Custom/Calibrate option.  You may save a new calibration and the txt file will be created.

- The text file must have the following format to be read correctly (9 items across):

1|pTAM, Lband, highfield, 40mod|41|8|0.02|122|0.53127|6.90|24365

2|sample|42|5|5|125|0.5|7.00|25000

in which the fields are arranged in this order (separated by |)

index|name|Gaussian LW|Lorentzian A LW|Lorentzian B LW|A/B Distance|Oxygen slope|pKa|Phosphate slope